# Chapter 5 Second Trimester Prenatal Screening; Results from a Large Screening Program

John Sherwin, Ph.D., Bob Currier, Ph.D., Fred Lorey, Ph.D. Chief, Genetic Disease Laboratory, State of California, Berkeley, CA 94710

Introduction, History and General Description of the California Expanded AFP Screening Program.

The Genetic Disease Branch of the California Department of Health Services [GDB] began prenatal screening for neural tube defects through the measurement of maternal serum alpha-fetoprotein [AFP] in 1985.  Over 2.5 million women [2,621,849] were screened between 1985 and 1995.  The analyte panel was expanded in 1995 to include two additional markers in maternal serum, chorionic gonadotropin [hCG] and unconjugated estriol [uE3] in 1995.  With the addition of these markers, GDB also began a screening program for Down syndrome and trisomy 18.  From the beginning of the triple marker screening program to the end of 2001, over 2.3 million women [2,329,429] were screened.

The screening program was established under the legislative authority of the California Code of Regulations, Title 17, Division 1, Chapter 4, Subchapter 9.  As specified by the regulations, prenatal care providers are required to offer the screening program to women under their care between 15 and 20 weeks of gestation.  Women then sign a document indicating that they consent to the screening or they decline the screening.  Women who consent to screening have a sample of blood drawn and sent for analysis.  Individuals are charged a fee, which will be paid by insurance or by MediCal.

The analysis is performed at one of eight regional laboratories under contract with GDB.  Upon arrival at the laboratory, the specimen is accessioned, including a determination whether or not the specimen is adequate for analysis.  The analyses are performed on multiple AutoDelphia Instruments (Perkin-Elmer Life Sciences, Boston, MA) using time-resolved fluorometry with reagents supplied by the instrument manufacturer.

The regional laboratories act under the direction of the central Genetic Disease Laboratory [GDL], which is responsible for quality assurance.  Daily results are monitored using prepared internal quality control materials at 8 different concentrations, as well as by monitoring patient medians, analytical tray medians and periodic external proficiency testing.  Results that pass Q/A measures are released to the central computer for interpretation.

Interpretation of the laboratory results depends not only on the analytical results from the laboratory but also on a number of demographic factors, the chief of which is gestational age.  The typical values of all three analytes change during pregnancy.  In order to create a common scale, values are converted into multiples of the median [MoM] by dividing by the population median for the given day of gestation.  Thus, the overall median MoM should be 1.00.  This MoM is further adjusted to take into account body weight [as a surrogate for blood volume] and ethnicity.  For those women who are insulin-dependent diabetics, a further adjustment is required.  Screening in twin pregnancies also requires a special adjustment.

Screening for neural tube defects [and abdominal wall defects] is based on fixed cut-offs of the AFP MoM of 2.5 for singleton pregnancies and 4.5 for multiple gestations.  This choice of cut-off results in a positive rate of approximately 1.5% and a detection rate in excess of 80% for the most common of these defects: anencephaly, (open) spina bifida, omphalocele and gastroschisis.

Screening for Down syndrome is based on a risk estimate.  The woman’s age provides the a priori risk, which is adjusted based on the likelihood ratio of the analyte values in Down syndrome pregnancies compared to unaffected pregnancies.  The resulting risk estimate is consider positive if the risk is greater than or equal to 1:190 at midtrimester.  This choice of risk cut-off gives an initial screen positive rate of 7%-8%.  Approximately one-third of the initial screen positives have overestimated gestational ages, giving falsely positive results.  The detection rate of the screening program exceeds 60%.  Since the population parameters in affected twin pregnancies are unknown, screening in twin pregnancies is performed by adjusting the analyte MoMs to the corresponding levels in a singleton pregnancy, and applying the risk algorithm.

Screening for trisomy 18 is similarly based on risk.  The a priori risk is again based on age, approximately one-tenth the Down syndrome risk.  Unconjugated estriol [uE3] is a particularly important marker for trisomy 18, so samples in which the uE3 is considered invalid do not receive a risk estimate for trisomy 18.  Multiple gestations are also not screened for trisomy 18.  Further, affected fetuses are frequently subject to growth retardation, so no changes to the initial estimate of gestational age are permitted.

Positive results are called to the attention of the prenatal care provider by the staff of seven regional Expanded AFP Coordinator offices.  The demographic data and other data upon which the positive result is based are confirmed by the Coordinator.  Then the patient is offered a referral for diagnostic procedures to one of the State-approved Prenatal Diagnosis Centers [PDC].  There are currently 27 Comprehensive Prenatal Diagnostic Centers with 109 satellite offices throughout California.

At the PDC, the patient is offered genetic counseling, detailed ultrasound and amniocentesis for diagnosis if indicated.  The costs of these follow-up services are reimbursed by the Expanded AFP Screening Program to the Prenatal Diagnostic Centers from the fees collected for the screening.

This summary of the screening program provided the basis for the consideration of two significant points for evaluation and monitoring.

The evaluation of kit lots.

The central laboratory [GDL] is also responsible for evaluation of new lots of reagents.  New kit lots are compared with the existing kit lot prior to use.  This comparison is done using both quality control material as reference materials tested over a period of 4 days on 2 different instruments.  Kit lots must match within 3% or the kit lot is referred back to the manufacturer for review and if necessary reformulation.  Only kit lots that match within the 3% limits and exhibit acceptable precision of better than 5% are placed into use.  The use of reference materials prevents the phenomenon of kit to kit drift since all lots are referenced against a known material

[ See Table I for sample results of a kit lot evaluation.]  On occasion, an assay may have a new formulation.  If preliminary quality assurance testing shows that the difference between the new assay and the current one will exceed 10%, then it is necessary to perform parallel testing with the two assays in order that the new assay can be interpreted with appropriate medians, i.e. medians derived from that assay itself (Table 2).

It is also important to monitor the variation of the assay.  Larger CVs lead to blurring the distinction between affected and unaffected pregnancies.  The result is both an increased false positive rate and a decreased detection rate.

Monitoring medians and positive rates (Figures 1 and 2).

At the population level, the fundamental outcome of the screening program is the screen positive rate, the percent of women who are identified for follow-up.  Changes in this rate in either direction can be significant: too high, and too many women are subject to invasive procedures, to say nothing of the increased cost of follow-up; too low, and too few affected fetuses are identified.

Further, monitoring the screen positive rate can point to the need for new adjustments.  Early in the California Program’s experience of triple marker testing, we observed a significant variation in the screen positive rates among the regional laboratories.  Further analysis showed that the median estriol varied with time from blood collection.  As a consequence, the Program instituted adjustments for small transit times and a policy of declaring estriol invalid when assayed more than eight days from blood collection.

Monitoring population medians is one tool to help identify causes of variation of screen positive rates.  One significant source of variation is ethnicity.  We now adjust all three analytes for ethnicity—each in different ways—based on the observed medians by ethnic group.  It is important both to collect ethnicity information in the population and to monitor differences in medians in subgroups to provide the basis for an adjustment.

A second major adjustment is for maternal weight.  Here, it is necessary to group the population in appropriately sized groups for the comparison.  In many cases, deciles will suffice.  An appropriate function [logarithmic or reciprocal] of the median analyte MoM in each group is regressed against the mean weight.  The result gives a functional dependence that can be applied to the population generally.

Other factors from smaller segments of the population that lead to adjustments within the California program include diabetes and twins.  In other programs there may be adjustments for smoking, previous history, number of pregnancies and fertility assistance.

The median MoMs represent the center of the population distribution, but the screening cut-offs are far out in the tails of the distribution.  Consequently, small changes in the median MoMs can be associated with large changes in the screen positive rates.  Screening for trisomy 18 is particularly sensitive to this phenomenon.

Recommendations

Alpha-fetoprotein [AFP] Assay

The coefficient of variation should not exceed 5%.

The accuracy should be within 3% from lot to lot.

Chorionic Gonadotropin [hCG] Assay

The coefficient of variation should not exceed 5%.

The accuracy should be within 3% from lot to lot.

Unconjugated estriol [uE3]

The coefficient of variation should not exceed 7%

The accuracy should be within 5% from lot to lot.

Since uE3 is not stable on storage at room temperature, programs should monitor time from specimen collection to analysis and reject specimens that are old enough to exhibit deterioration

# The Screening Program

Since all three analytes exhibit variation by race, data should be collected in order to adjust MoMs appropriately, if the required correction is greater than 10%.

Repeat testing of initial positive results should not be performed because of the correlation of successive assays and the phenomenon of regression to the mean.

Programs doing second trimester screening should consider adding inhibin A and/or invasive trophoblast antigen [ITA]

Table 1. Validation of a new kit lot.

| Interpretation | Start Date | Change of |  | Kit Start | Number of | Gestation | and Weight | and Ethnicity | Number of | Gestation | and Weight | and Ethnicity |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Version |  | Medians | AFP Kit Lot | Date | Samples | Median | log SD | Median | log SD | Median | log SD | Samples | Median | log SD | Median | log SD | Median | log SD |
| 6.00 | 07/29/95 | X | 111111 | 7/29/1995 | 4399 | 1.03 | 0.1837 | 1.02 | 0.1745 | 1.01 | 0.1724 | 2419 | 1.02 | 0.1570 | 1.01 | 0.1495 | 1.00 | 0.1518 |
| 6.00 |  |  | 222222 | 8/8/1995 | 51326 | 1.07 | 0.1830 | 1.05 | 0.1711 | 1.04 | 0.1722 | 28398 | 1.05 | 0.1601 | 1.05 | 0.1479 | 1.04 | 0.1472 |
| 6.01 | 11/06/95 |  |  |  | 25430 | 1.07 | 0.1793 | 1.05 | 0.1687 | 1.05 | 0.1674 | 15527 | 1.06 | 0.1578 | 1.06 | 0.1464 | 1.05 | 0.1479 |
| 6.01 |  |  | 222222-plate | 12/21/1995 | 15458 | 1.03 | 0.1798 | 1.02 | 0.1720 | 1.01 | 0.1713 | 9395 | 1.01 | 0.1575 | 1.01 | 0.1487 | 1.00 | 0.1472 |
| 6.10 | 01/18/96 |  |  |  | 75787 | 1.03 | 0.1817 | 1.02 | 0.1720 | 1.01 | 0.1706 | 47410 | 1.02 | 0.1575 | 1.02 | 0.1494 | 1.01 | 0.1472 |
| 6.10 |  |  | 222222-plate | 5/22/1996 | 12269 | 1.04 | 0.1797 | 1.02 | 0.1692 | 1.01 | 0.1685 | 7771 | 1.02 | 0.1664 | 1.02 | 0.1531 | 1.01 | 0.1524 |
| 6.10 |  |  | 333333-std | 6/12/1996 | 6442 | 1.04 | 0.1849 | 1.03 | 0.1791 | 1.03 | 0.1778 | 3976 | 1.03 | 0.1647 | 1.03 | 0.1537 | 1.02 | 0.1552 |
| 6.10 |  |  | 334333-tracer | 6/22/1996 | 9411 | 1.04 | 0.1817 | 1.03 | 0.1755 | 1.02 | 0.1773 | 5977 | 1.02 | 0.1671 | 1.01 | 0.1559 | 1.01 | 0.1552 |
| 6.11 | 07/10/96 |  |  |  | 33670 | 1.04 | 0.1868 | 1.03 | 0.1760 | 1.02 | 0.1740 | 22052 | 1.03 | 0.1610 | 1.03 | 0.1486 | 1.02 | 0.1501 |
| 6.11 |  |  | 334333-plate | 9/9/1996 | 24498 | 1.05 | 0.1849 | 1.03 | 0.1742 | 1.03 | 0.1728 | 16685 | 1.04 | 0.1630 | 1.04 | 0.1501 | 1.03 | 0.1516 |
| 6.11 |  |  | 334343-tracer | 10/23/1996 | 57534 | 1.02 | 0.1823 | 1.01 | 0.1720 | 1.00 | 0.1731 | 40704 | 1.01 | 0.1599 | 1.01 | 0.1487 | 1.00 | 0.1502 |
| 7.00 | 02/07/97 |  |  |  | 30887 | 1.02 | 0.1810 | 1.01 | 0.1745 | 1.00 | 0.1731 | 22911 | 1.00 | 0.1601 | 1.00 | 0.1495 | 0.99 | 0.1503 |
| 7.00 |  |  | 334443-plate | 4/3/1997 | 7418 | 1.00 | 0.1777 | 0.99 | 0.1671 | 0.99 | 0.1650 | 5539 | 0.99 | 0.1580 | 0.99 | 0.1480 | 0.98 | 0.1465 |
| 7.01 | 04/16/97 |  |  |  | 62491 | 1.01 | 0.1823 | 0.99 | 0.1724 | 0.99 | 0.1710 | 47064 | 0.99 | 0.1601 | 0.99 | 0.1488 | 0.98 | 0.1473 |
| 7.10 | 08/12/97 |  |  |  | 26368 | 1.01 | 0.1843 | 1.00 | 0.1738 | 1.00 | 0.1724 | 21632 | 0.99 | 0.1608 | 0.99 | 0.1488 | 0.99 | 0.1480 |
| 7.10 |  |  | 334543-plate | 10/2/1997 | 68304 | 1.05 | 0.1843 | 1.05 | 0.1748 | 1.04 | 0.1735 | 56877 | 1.04 | 0.1624 | 1.04 | 0.1493 | 1.03 | 0.1501 |
| 7.10 |  |  | 444544-std | 2/19/1998 | 13393 | 1.06 | 0.1817 | 1.06 | 0.1700 | 1.05 | 0.1680 | 11527 | 1.04 | 0.1587 | 1.05 | 0.1464 | 1.04 | 0.1457 |
| 7.10 |  |  | 444654-plate | 3/16/1998 | 31417 | 1.03 | 0.1791 | 1.03 | 0.1695 | 1.02 | 0.1681 | 26830 | 1.01 | 0.1561 | 1.02 | 0.1465 | 1.01 | 0.1458 |
| 7.10 |  |  | 445664-tracer | 5/14/1998 | 28091 | 1.05 | 0.1792 | 1.05 | 0.1680 | 1.05 | 0.1667 | 23985 | 1.04 | 0.1587 | 1.05 | 0.1464 | 1.04 | 0.1479 |
| 7.10 |  |  | 683741 | 7/9/1998 | 3222 | 1.07 | 0.1799 | 1.08 | 0.1719 | 1.06 | 0.1719 | 2678 | 1.09 | 0.1602 | 1.10 | 0.1464 | 1.09 | 0.1450 |

# Table 2 The effect of reformulation of the unconjugated estriol assay

| Week of Gestation | 711530 | 717541 | Percent Increase |
| --- | --- | --- | --- |
|  | LMP | U/S | LMP | U/S | LMP | U/S |
| 15 | 0.84 | 0.85 | 0.89 | 0.91 | 5.95% | 7.06% |
| 16 | 1.04 | 1.06 | 1.13 | 1.15 | 8.65% | 8.49% |
| 17 | 1.32 | 1.33 | 1.45 | 1.46 | 9.85% | 9.77% |
| 18 | 1.61 | 1.61 | 1.78 | 1.78 | 10.56% | 10.56% |
| 19 | 1.9 | 1.89 | 2.14 | 2.17 | 12.63% | 14.81% |
| 20 | 2.19 | 2.13 | 2.49 | 2.44 | 13.70% | 14.55% |

# Figure 1.  Monitoring assay medians and Down Syndrome screen positive rates  Data are shown for last menstrual period dating.

# Figure 2.  Monitoring assay medians and Down Syndrome screen positive rates  Data are shown for ultrasound dating.

# References

Laboratory assay

1. Wald N.J., Cuckle H.S., The quality control of alpha-fetoprotein reagents and assay for the antenatal axreening and diagnosis of open neural-tube defects.  Report of a workshop sponsored by the National Institute of Child Health and Human Development.  Clin. Chem. Acta 1980; 105:9—24.

2

3

# Medians, Adjustments and Screening

4. Wald N.J., Kennard A., et al., Antenatal screening for Down’s syndrome, J. Med. Screening, 1997: 4 (4):181—246.  A review article.  Section 3 on second-trimester serum markers contains a bibliography of 149 items.

5. Wald N.J., Brock D.J.H., Bonnar J., Prenatal diagnosis of spina bifida and anencephaly by maternal serum alpha-fetoprotein measurement.  Lancet 1974; i:765—767.

6. Cuckle H.S., Wald N.J., Thompson S.G., Estimating a woman’s risk of having a pregnancy with Down’s syndrome using her age and serum alpha-fetoprotein level., Br. J. Obstet. Gynaecol., 1987; 94:387—402.

7. Haddow J.E., Palomaki G.E., et al., Prenatal screening for Down’s syndrome with use of maternal serum markers, N Engl J Med 1992; 327:588—593.

8. Palomaki G.E., et al., Risk-based screening for trisomy 18 using alpha-fetoprotein, unconjugated oestriol and human chorionic gonadotropin, Prenatal Diagnosis 1995; 15:713—723.

# Footnotes

This was formerly known as hyperglycosylated hCG [HhCG].

# Endnotes

Kaiser Medical Group Berkeley, Kaiser Medical Group Los Angeles, Western Clinical Laboratory, Allied Clinical Laboratory,